| Steel Septum Magnets for the LHC Injection system |
| --- |
| Technical Inspection Report |
| Magnet Serial Number: Lhc-msiВ-01 |
| Prepared by : Name IHEP A.Ivanenko N.Yarygin V.Noskov E.Gudkova | Checked by : Name IHEP Ludmirsky@mx.ihep.su | Approved by: Name CERN Manfred.Sassowsky@cern.ch |

Table of Contents

1. Half Cores	3

1.2    Septum Core	3

1.2    Coil core	4

2.	Excitation coil	5

2.1    Mechanical Tolerances	5

2.2    Flow and Pressure tests	5

2.3    electrical resistance tests	6

2.4    Insulation tests	6

3. Assembled Magnet	7

3.1    Mechanical Tolerances	7

3.2    leak and water flow tests	8

3.3    Electrical Resistance tests	8

3.4    Test run	8

4. Magnetic Measurements	9

4.1    System of axes MSI	9

4.2    Magnetisation curve	9

4.3    Fied map in the gap	10

4.4    field in the holes	13

# 1. Half Cores

## 1.2 Septum Core

# Technical Details

| Serial Number | LHC-MSIB-01 |
| --- | --- |
| Steel stock used for core | 50080/2,9,11 |
| Date of Inspection | 20.06.2000 |
| Inspector | V.Noskov |
| Stacking Factor | 0.98 |

Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Length | 4000 | 3998 |
| Width | 714 | 714 |
| Height | 195.2 | 195.2 |
| Straightness 1) | 0.3 | 0.3 |
| 2) | 0.3 | 0.3 |
| Perpendicularity | 0.5 | 0.5 |

# Weight [kg]

|  | Nominal | Measured |
| --- | --- | --- |
| Laminations | 3363.74 | 3411.66 |
| End-plates | 53.22 | 52.00 |
| Angular plates | 273.04 | 269.4 |
| Total weight | 3690 | 3733.06 |

1.2 Coil core

# Technical Details

| Serial Number | LHC-MSIB-01 |
| --- | --- |
| Steel stock used for core | 50080/9,11 |
| Date of Inspection | 20.06.2000 |
| Inspector | V.Noskov |
| Stacking Factor | 0.98 |

Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Length | 3650 | 3648 |
| Width | 714 | 714 |
| Height | 308.8 | 308.8 |
| Straightness 1) | 0.3 | 0.3 |
| 2) | 0.3 | 0.3 |
| Perpendicularity | 0.5 | 0.5 |

# Weight [kg]

|  | Nominal | Measured |
| --- | --- | --- |
| Laminations | 5039.68 | 5175.65 |
| End-plates | 88.02 | 88.00 |
| Angular plates | 312.3 | 308.40 |
| T0otal weight | 5440 | 5572.05 |

# Excitation coil

# Technical Details

| Serial Number | LHC-MSIB-01 |
| --- | --- |
| Date of Inspection | 25.07.2000 |
| Inspector | N.Yarygin |

## 2.1 Mechanical Tolerances

# Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Length | 3900 | 3902 |
| Width | 373.6 | 373.9 |
| Height | 100 | 100 |
| Connection Side | 109* | 103 |
| Opposite Side | 109* | 102 |
| Straightness | 0.5/1000 | 0.45/1000 |

## 2.2 Flow and Pressure tests

# Flow Test

|  | Unit | Nominal | Measured |
| --- | --- | --- | --- |
| Pressure drop | Bar | 5 | 5 |
| Flow layer 1 | L/min | 1.97 | 2.2 |
| Flow layer 2 | L/min | 1.97 | 2.2 |
| Flow layer 3 | L/min | 1.97 | 2.2 |
| Flow layer 4 | L/min | 1.97 | 2.2 |
| Flow layer 5 | L/min | 1.97 | 2.2 |
| Flow layer 6 | L/min | 1.97 | 2.2 |
| Total flow | L/min | 11.8 | 13.2 |

# Pressure Test

|  | Nominal | Measured |
| --- | --- | --- |
| Pressure (bar) | 60 | 60 |
| Time (min) | 60 | 60 |
| Evidence of Leaks | No |

## 2.3 electrical resistance tests

# Electrical resistance

|  | Nominal | Measured |
| --- | --- | --- |
| Temperature (°C) | 20 | 20 |
| Resistance (mohm) | 16.4 | 16.2 |

## 2.4 Insulation tests

# Test Conditions, water immersion

|  | Nominal | Measured |
| --- | --- | --- |
| Temperature (°C) | 20 | 20 |
| Time (min.) | 60 | 360 |
| A.C. Voltage (1 minute) | 4.4 kV 50 Hz | 4.4 kV 50 Hz |

Test Results

|  | Nominal | Measured |
| --- | --- | --- |
| Resistance before | (108 Ω | 2 108 ( |
| Resistance after | (108 Ω | 2 108 ( |

# Inter turn resistance after immersion

|  | Nominal | Measured |
| --- | --- | --- |
| applied Voltage | 750V rms | 750 V rms |
| Layers 1-2 | (108 Ω | 2 108 ( |
| Layers 2-3 | (108 Ω | 2 108 ( |
| Layers 3-4 | (108 Ω | 2 108 ( |
| Layers 4-5 | (108 Ω | 2 108 ( |
| Layers 5-6 | (108 Ω | 2 108 ( |

# 3. Assembled Magnet

# Technical Details

| Serial Number | LHC-MSIB-01 |
| --- | --- |
| Date of Inspection | 11.08.2000 |
| Inspector | E.Gudkova |

## 3.1 Mechanical Tolerances

# Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Transversal offset | 0.3 | 0.3 |
| Clearance | 0.05 | 0.05 |
| Vertical Straightness | 0.3 | 0.3 |
| Horizontal Straightness | 0.15 | 0.17 |
| Twist | 0.75 mrad | ( 0.6 mrad |

## 3.2 leak and water flow tests

# Flow Test

|  | Unit | Nominal | Measured |
| --- | --- | --- | --- |
| Pressure drop | Bar | 5 | 5 |
| Total flow | L/min | 11.8 | 12.9 |

# Pressure Test

|  | Nominal | Measured |
| --- | --- | --- |
| Pressure (bar) | 60 | 60 |
| Time (min) | 60 | 60 |
| Evidence of Leaks | No |

## 3.3 Electrical Resistance tests

# Electrical resistance at 20°C

|  | Nominal | Measured |
| --- | --- | --- |
| Temperature (°C) | 20 | 21 |
| Resistance (mohm) | 16.4 | 16.3 |

Electrical resistance between coil and ground

|  | Nominal | Measured |
| --- | --- | --- |
| Voltage (kV d.c.) | 20 | 2.0 |
| Temperature (°C) | 20 | 21 |
| Resistance (Ω | Min. 108 | (108 Ω |

## 3.4 Test run

# Cooling performance during test run at nominal current

|  | Checked Cutoff Cut-in |
| --- | --- |
| Functioning of thermal switch 1 | 63 0 C / 44 0 C |
| Functioning of thermal switch 2 | 65 0 C / 46 0 C |
| Functioning of thermal switch 3 | 64 0 C / 43 0 C |
| Functioning of thermal switch 4 | 64 0 C / 45 0 C |
| Functioning of thermal switch 5 | 65 0 C / 44 0 C |
| Functioning of thermal switch 6 | 66 0 C / 45 0 C |

# 4. Magnetic Measurements

# Technical Details

| Serial Number | LHC-MSIB-01 |
| --- | --- |
| Date of Inspection | 20.03.2001 |
| Inspector | A.Ivanenko |

## 4.1 System of axes MSI

## 4.2 Magnetisation curve

Flux density ( inside the magnet and integrated flux density (stretched wire method) over the total length (Bdl as a function of the excitation current.

| EXCITING CURRENT, Ampere | INTEGRAL, VOLT | B(L, Tm | Magnetic Length L, m | NMR, T |
| --- | --- | --- | --- | --- |
| 50.0 | 0.02466 | 0.21804 | 3.62077 | 0.06022 |
| 100.0 | 0.04963 | 0.43883 | 3.68917 | 0.11895 |
| 150.0 | 0.07500 | 0.66315 | 3.69854 | 0.17930 |
| 200.0 | 0.10033 | 0.88711 | 3.70712 | 0.23930 |
| 250.0 | 0.12581 | 1.11241 | 3.71918 | 0.29910 |
| 300.0 | 0.15089 | 1.33416 | 3.71582 | 0.35905 |
| 350.0 | 0.17586 | 1.55495 | 3.71819 | 0.41820 |
| 400.0 | 0.20136 | 1.78042 | 3.72784 | 0.47760 |
| 450.0 | 0.22558 | 1.99457 | 3.71567 | 0.53680 |
| 500.0 | 0.25061 | 2.21589 | 3.71606 | 0.59630 |
| 550.0 | 0.27564 | 2.43720 | 3.72035 | 0.65510 |
| 600.0 | 0.30022 | 2.65454 | 3.71758 | 0.71405 |
| 650.0 | 0.32551 | 2.87815 | 3.72383 | 0.77290 |
| 700.0 | 0.35013 | 3.09584 | 3.72409 | 0.83130 |
| 750.0 | 0.37476 | 3.31362 | 3.72400 | 0.88980 |
| 800.0 | 0.39912 | 3.52901 | 3.72258 | 0.94800 |
| 850.0 | 0.42298 | 3.73997 | 3.71767 | 1.00600 |
| 900.0 | 0.44765 | 3.95811 | 3.72212 | 1.06340 |
| 950.0 | 0.47083 | 4.16306 | 3.71835 | 1.11960 |
| 1000.0 | 0.49347 | 4.36325 | 3.71498 | 1.17450 |
| 1050.0 | 0.51493 | 4.55299 | 3.72390 | 1.22264 |
| 1100.0 | 0.53424 | 4.72373 | 3.71363 | 1.27200 |
| 1150.0 | 0.55222 | 4.88271 | 3.69454 | 1.32160 |

## 4.3 Fied map in the gap

|  |  | 5 mm towards pole |  |  | Current 950 A |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Int (B dl) [Tm] |  |  |  |  |  |  |  |
|  | End | Center | All | Homogeneity [%] |  | B_ave_center [T] |  | l_eff [m] |
| X [mm] |  |  |  |  |  |  |  |  |
| 150 | 0,231996 | 3,952221 | 4,184217 | 0,674744 |  | 1,130038 |  | 3,702721 |
| 145 | 0,233771 | 3,926811 | 4,160582 | 0,106057 |  | 1,122798 |  | 3,705546 |
| 140 | 0,234888 | 3,922422 | 4,15731 | 0,027339 |  | 1,121532 |  | 3,706813 |
| 135 | 0,236158 | 3,922293 | 4,158451 | 0,054794 |  | 1,121498 |  | 3,707944 |
| 130 | 0,236786 | 3,922572 | 4,159358 | 0,076627 |  | 1,121564 |  | 3,708534 |
| 125 | 0,237649 | 3,923818 | 4,161467 | 0,127367 |  | 1,121877 |  | 3,709378 |
| 115 | 0,240503 | 3,92396 | 4,164463 | 0,199447 |  | 1,12191 |  | 3,711941 |
| 105 | 0,241533 | 3,922454 | 4,163987 | 0,187987 |  | 1,121507 |  | 3,71285 |
| 95 | 0,24189 | 3,922601 | 4,164491 | 0,200117 |  | 1,121526 |  | 3,713237 |
| 85 | 0,242195 | 3,920328 | 4,162524 | 0,152785 |  | 1,120905 |  | 3,713539 |
| 75 | 0,243119 | 3,918654 | 4,161774 | 0,134735 |  | 1,120431 |  | 3,71444 |
| 65 | 0,243521 | 3,917672 | 4,161193 | 0,120758 |  | 1,120124 |  | 3,71494 |
| 55 | 0,243493 | 3,91543 | 4,158923 | 0,066145 |  | 1,119501 |  | 3,714977 |
| 45 | 0,243371 | 3,915223 | 4,158594 | 0,058234 |  | 1,119411 |  | 3,714985 |
| 35 | 0,243251 | 3,914124 | 4,157375 | 0,02891 |  | 1,119093 |  | 3,714951 |
| 25 | 0,243841 | 3,915481 | 4,159322 | 0,075739 |  | 1,119483 |  | 3,715395 |
| 20 | 0,243511 | 3,915144 | 4,158655 | 0,059712 |  | 1,119374 |  | 3,715162 |
| 15 | 0,243179 | 3,915361 | 4,158539 | 0,05692 |  | 1,119433 |  | 3,714862 |
| 10 | 0,24215 | 3,91547 | 4,157619 | 0,034776 |  | 1,119474 |  | 3,713905 |
| 5 | 0,241825 | 3,914658 | 4,156483 | 0,007447 |  | 1,11926 |  | 3,713599 |
| 0 | 0,241686 | 3,914792 | 4,156477 | 0,007303 |  | 1,119282 |  | 3,713521 |
| -5 | 0,241607 | 3,914668 | 4,156275 | 0,002434 |  | 1,119234 |  | 3,713499 |
| -10 | 0,241428 | 3,915115 | 4,156543 | 0,008878 |  | 1,119371 |  | 3,713285 |
| -15 | 0,240859 | 3,915032 | 4,155891 | -0,00681 |  | 1,119336 |  | 3,712819 |
| -20 | 0,239837 | 3,915909 | 4,155746 | -0,01029 |  | 1,119615 |  | 3,711762 |
| -25 | 0,23859 | 3,9172 | 4,155789 | -0,00925 |  | 1,120005 |  | 3,710509 |
| -30 | 0,235596 | 3,915457 | 4,151053 | -0,12322 |  | 1,119509 |  | 3,707922 |
| -35 | 0,233571 | 3,92012 | 4,15369 | -0,05975 |  | 1,120788 |  | 3,706045 |
| -40 | 0,232165 | 3,949425 | 4,18159 | 0,611528 |  | 1,129126 |  | 3,703386 |

|  |  | mid-plane |  | Current 950 A |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Int (B dl) [Tm] |  |  |  |  |  |  |  |
|  | End | Center | All | Homogeneity [%] |  | B_ave_center [T] |  | l_eff [m] |
| X [mm] |  |  |  |  |  |  |  |  |
| 150 | 0,229596 | 3,957039 | 4,186635 | 0,732913 |  | 1,131489 |  | 3,700111 |
| 145 | 0,23239 | 3,934205 | 4,166595 | 0,250735 |  | 1,124873 |  | 3,704057 |
| 140 | 0,234131 | 3,925457 | 4,159588 | 0,082144 |  | 1,122362 |  | 3,706101 |
| 135 | 0,235857 | 3,922826 | 4,158684 | 0,060389 |  | 1,121597 |  | 3,707824 |
| 130 | 0,237062 | 3,922895 | 4,159957 | 0,09103 |  | 1,121624 |  | 3,708869 |
| 125 | 0,238182 | 3,923411 | 4,161593 | 0,130392 |  | 1,121736 |  | 3,709958 |
| 115 | 0,240855 | 3,923675 | 4,164531 | 0,201076 |  | 1,121795 |  | 3,71238 |
| 105 | 0,241872 | 3,922338 | 4,16421 | 0,193365 |  | 1,121438 |  | 3,713277 |
| 95 | 0,242438 | 3,922631 | 4,165068 | 0,214007 |  | 1,121504 |  | 3,713823 |
| 85 | 0,243116 | 3,920812 | 4,163928 | 0,18658 |  | 1,121019 |  | 3,714413 |
| 75 | 0,243662 | 3,918798 | 4,16246 | 0,151256 |  | 1,120439 |  | 3,715026 |
| 65 | 0,243866 | 3,917076 | 4,160941 | 0,114712 |  | 1,119917 |  | 3,7154 |
| 55 | 0,244017 | 3,915564 | 4,159581 | 0,081979 |  | 1,119506 |  | 3,715549 |
| 45 | 0,244198 | 3,91497 | 4,159168 | 0,072038 |  | 1,119304 |  | 3,715852 |
| 35 | 0,24367 | 3,91435 | 4,15802 | 0,04443 |  | 1,11914 |  | 3,715372 |
| 25 | 0,244019 | 3,914676 | 4,158695 | 0,06066 |  | 1,119214 |  | 3,715726 |
| 20 | 0,243704 | 3,914631 | 4,158335 | 0,051992 |  | 1,119193 |  | 3,715476 |
| 15 | 0,243453 | 3,915076 | 4,158528 | 0,056651 |  | 1,119316 |  | 3,715241 |
| 10 | 0,242542 | 3,914727 | 4,157269 | 0,026354 |  | 1,11922 |  | 3,714433 |
| 5 | 0,242329 | 3,914323 | 4,156652 | 0,011518 |  | 1,119124 |  | 3,714203 |
| 0 | 0,241955 | 3,914218 | 4,156174 | 0 |  | 1,11908 |  | 3,713921 |
| -5 | 0,24161 | 3,914368 | 4,155978 | -0,00471 |  | 1,119109 |  | 3,71365 |
| -10 | 0,24111 | 3,913727 | 4,154836 | -0,03217 |  | 1,118939 |  | 3,713195 |
| -15 | 0,240321 | 3,913813 | 4,154134 | -0,04907 |  | 1,118957 |  | 3,712504 |
| -20 | 0,239058 | 3,914478 | 4,153536 | -0,06347 |  | 1,119177 |  | 3,711242 |
| -25 | 0,237701 | 3,916322 | 4,154023 | -0,05175 |  | 1,11972 |  | 3,709875 |
| -30 | 0,23492 | 3,917422 | 4,152342 | -0,0922 |  | 1,120053 |  | 3,707271 |
| -35 | 0,232886 | 3,923895 | 4,156781 | 0,014611 |  | 1,121888 |  | 3,705166 |
| -40 | 0,230557 | 3,9458 | 4,176357 | 0,485625 |  | 1,128186 |  | 3,701834 |

|  |  | 5 mm towards septum |  |  | Current 950 A |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Int (B dl) [Tm] |  |  |  |  |  |  |  |
|  | End | Center | All | Homogeneity [%] |  | B_ave_center [T] |  | l_eff [m] |
| X [mm] |  |  |  |  |  |  |  |  |
| 150 | 0,225323 | 3,922264 | 4,147587 | -0,20661 |  | 1,121515 |  | 3,698199 |
| 145 | 0,229292 | 3,925214 | 4,154505 | -0,04014 |  | 1,122286 |  | 3,701826 |
| 140 | 0,231674 | 3,922876 | 4,154551 | -0,03905 |  | 1,121616 |  | 3,704076 |
| 135 | 0,233819 | 3,923048 | 4,156867 | 0,016691 |  | 1,121663 |  | 3,705985 |
| 130 | 0,235391 | 3,922612 | 4,158002 | 0,043991 |  | 1,121523 |  | 3,70746 |
| 125 | 0,236672 | 3,923806 | 4,160478 | 0,103565 |  | 1,121859 |  | 3,708556 |
| 115 | 0,239404 | 3,923844 | 4,163248 | 0,170207 |  | 1,121857 |  | 3,711034 |
| 105 | 0,240502 | 3,922753 | 4,163255 | 0,170381 |  | 1,121542 |  | 3,712081 |
| 95 | 0,241154 | 3,92297 | 4,164124 | 0,191284 |  | 1,121617 |  | 3,712606 |
| 85 | 0,242027 | 3,920553 | 4,16258 | 0,154138 |  | 1,120916 |  | 3,713552 |
| 75 | 0,242479 | 3,918606 | 4,161085 | 0,118169 |  | 1,120363 |  | 3,714051 |
| 65 | 0,242644 | 3,91778 | 4,160423 | 0,102248 |  | 1,12014 |  | 3,7142 |
| 55 | 0,24289 | 3,91573 | 4,15862 | 0,058868 |  | 1,119541 |  | 3,714577 |
| 45 | 0,243172 | 3,91537 | 4,158542 | 0,056983 |  | 1,11943 |  | 3,714875 |
| 35 | 0,242583 | 3,91456 | 4,157143 | 0,02333 |  | 1,119175 |  | 3,714471 |
| 25 | 0,242783 | 3,914975 | 4,157758 | 0,038122 |  | 1,119328 |  | 3,714514 |
| 20 | 0,242458 | 3,914859 | 4,157317 | 0,027508 |  | 1,119278 |  | 3,714285 |
| 15 | 0,242233 | 3,915106 | 4,157339 | 0,028035 |  | 1,119346 |  | 3,714079 |
| 10 | 0,24141 | 3,915407 | 4,156817 | 0,01548 |  | 1,119435 |  | 3,713319 |
| 5 | 0,241248 | 3,914696 | 4,155944 | -0,00553 |  | 1,119224 |  | 3,713238 |
| 0 | 0,240779 | 3,914682 | 4,155461 | -0,01715 |  | 1,119225 |  | 3,712801 |
| -5 | 0,24023 | 3,914207 | 4,154437 | -0,04178 |  | 1,119091 |  | 3,712333 |
| -10 | 0,239655 | 3,914387 | 4,154042 | -0,05129 |  | 1,119144 |  | 3,711804 |
| -15 | 0,238685 | 3,914133 | 4,152818 | -0,08073 |  | 1,119072 |  | 3,71095 |
| -20 | 0,237262 | 3,914585 | 4,151847 | -0,10411 |  | 1,119211 |  | 3,709618 |
| -25 | 0,235709 | 3,915089 | 4,150798 | -0,12934 |  | 1,119389 |  | 3,708092 |
| -30 | 0,232758 | 3,91327 | 4,146028 | -0,24411 |  | 1,118868 |  | 3,705556 |
| -35 | 0,230437 | 3,9138 | 4,144237 | -0,2872 |  | 1,119042 |  | 3,70338 |
| -40 | 0,227051 | 3,913908 | 4,140958 | -0,36609 |  | 1,119109 |  | 3,700227 |

## 4.4 field in the holes

1. Hall probe (radial),        R = 19 mm (Initial angular position equal     00

2. Hall probe (tangential)   R = 19 mm ( Initial angular position equal 2700

| S E P T U M H O L E | B E A M H O L E |
| --- | --- |

| Position degree of circle | Hall probe (radial), Gauss | Hall probe (tangential), Gauss | Absolute Value, Gauss |  | Hall probe (radial), Gauss | Hall probe (tangential), Gauss | Absolute Value, Gauss |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 8.85 | - 2.17 | 9.10 |  | - 6.80 | - 4.80 | 8.38 |
| 30 | 8.00 | - 4.90 | 9.38 |  | - 7.80 | - 2.60 | 8.22 |
| 60 | 6.40 | - 8.90 | 10.96 |  | - 8.90 | 0.25 | 8.91 |
| 90 | 3.16 | - 12.70 | 13.09 |  | - 9.30 | 5.02 | 10.57 |
| 120 | - 4.80 | - 15.40 | 16.13 |  | - 5.90 | 11.24 | 12.69 |
| 150 | - 16.10 | - 10.60 | 13.45 |  | 4.36 | 12.60 | 13.33 |
| 180 | - 19.86 | 0.00 | 19.86 |  | 10.90 | 6.43 | 12.65 |
| 210 | - 18.00 | 13.50 | 22.50 |  | 11.50 | - 0.42 | 11.51 |
| 240 | - 1.55 | 21.10 | 21.15 |  | 9.83 | - 6.10 | 11.57 |
| 270 | 10.70 | 11.40 | 15.63 |  | 5.20 | - 10.60 | 11.81 |
| 300 | 11.60 | 4.40 | 12.41 |  | - 1.60 | - 10.70 | 10.82 |
| 330 | 10.10 | 0.30 | 10.10 |  | - 5.50 | - 7.80 | 9.54 |
| 360 | 8.90 | - 2.17 | 9.16 |  | - 6.80 | - 4.90 | 8.38 |