| Steel Septum Magnets for the LHC Injection system |
| --- |
| Technical Inspection Report |
| Magnet Serial Number: Lhc-msiA-01 |
| Prepared by : Name IHEP A.Ivanenko N.Yarygin V.Noskov E.Gudkova | Checked by : Name IHEP Ludmirsky@mx.ihep.su | Approved by: Name CERN Delphine.Gerard@cern.ch |

Table of Contents

1. Half Cores	3

1.2    Septum Core	3

1.2    Coil core	4

2.	Excitation coil	5

2.1    Mechanical Tolerances	5

2.2    Flow and Pressure tests	5

2.3    electrical resistance tests	6

2.4    Insulation tests	6

3. Assembled Magnet	7

3.1    Mechanical Tolerances	7

3.2    leak and water flow tests	8

3.3    Electrical Resistance tests	8

3.4    Test run	8

4. Magnetic Measurements	9

4.1    System of axes MSI	9

4.2    Magnetisation curve	9

4.3    Fied map in the gap	9

4.4    field in the holes	13

# 1. Half Cores

## 1.2 Septum Core

# Technical Details

| Serial Number | LHC-MSIA-01 |
| --- | --- |
| Steel stock used for core | 50074/16 |
| Date of Inspection | 27.07.2001 |
| Inspector | V.Noskov |
| Stacking Factor | 0.98 |

Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Length | 4000 | 3999 |
| Width | 714 | 714,6 |
| Height | 195.2 | 195.7 |
| Straightness 1) | 0.3 | 0.3 |
| 2) | 0.3 | 0.3 |
| Perpendicularity | 0.5 | 0.61 |

# Weight [kg]

|  | Nominal | Measured |
| --- | --- | --- |
| Laminations | 3363.74 | 3407.18 |
| End-plates | 53.22 | 52.65 |
| Angular plates | 273.04 | 268.0 |
| Total weight | 3690 | 3727.83 |

1.2 Coil core

# Technical Details

| Serial Number | LHC-MSIA-01 |
| --- | --- |
| Steel stock used for core | 50080/18,21 |
| Date of Inspection | 20.09.2001 |
| Inspector | V.Noskov |
| Stacking Factor | 0.98 |

Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Length | 3650 | 3649 |
| Width | 714 | 714,8 |
| Height | 308.8 | 308.7 |
| Straightness 1) | 0.3 | 0.3 |
| 2) | 0.3 | 0.3 |
| Perpendicularity | 0.5 | 0.63 |

# Weight [kg]

|  | Nominal | Measured |
| --- | --- | --- |
| Laminations | 5039.68 | 5194.72 |
| End-plates | 88.02 | 88.65 |
| Angular plates | 312.3 | 306.50 |
| T0otal weight | 5440 | 5589.87 |

# Excitation coil

# Technical Details

| Serial Number | LHC-MSIA-01 |
| --- | --- |
| Date of Inspection | 16.10.2001 |
| Inspector | N.Yarygin |

## 2.1 Mechanical Tolerances

# Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Length | 3900 | 3901 |
| Width | 373.6 | 374.1 |
| Height | 100 | 100 |
| Connection Side | 109* | 103 |
| Opposite Side | 109* | 102 |
| Straightness | 0.5/1000 | 0.51/1000 |

## 2.2 Flow and Pressure tests

# Flow Test

|  | Unit | Nominal | Measured |
| --- | --- | --- | --- |
| Pressure drop | Bar | 5 | 5 |
| Flow layer 1 | L/min | 1.97 | 2.0 |
| Flow layer 2 | L/min | 1.97 | 2.0 |
| Flow layer 3 | L/min | 1.97 | 2.0 |
| Flow layer 4 | L/min | 1.97 | 2.0 |
| Total flow | L/min | 7.88 | 8.0 |

# Pressure Test

|  | Nominal | Measured |
| --- | --- | --- |
| Pressure (bar) | 60 | 60 |
| Time (min) | 60 | 60 |
| Evidence of Leaks | No |

## 2.3 electrical resistance tests

# Electrical resistance

|  | Nominal | Measured |
| --- | --- | --- |
| Temperature (°C) | 20 | 20 |
| Resistance (mohm) | 10.9 | 10.8 |

## 2.4 Insulation tests

# Test Conditions, water immersion

|  | Nominal | Measured |
| --- | --- | --- |
| Temperature (°C) | 20 | 20 |
| Time (min.) | 60 | 360 |
| A.C. Voltage (1 minute) | 4.4 kV 50 Hz | 4.4 kV 50 Hz |

Test Results

|  | Nominal | Measured |
| --- | --- | --- |
| Resistance before | (108 Ω | 2 108 ( |
| Resistance after | (108 Ω | 2 108 ( |

# Inter turn resistance after immersion

|  | Nominal | Measured |
| --- | --- | --- |
| applied Voltage | 750V rms | 750 V rms |
| Layers 1-2 | (108 Ω | 2 108 ( |
| Layers 2-3 | (108 Ω | 2 108 ( |
| Layers 3-4 | (108 Ω | 2 108 ( |

# 3. Assembled Magnet

# Technical Details

| Serial Number | LHC-MSIA-01 |
| --- | --- |
| Date of Inspection | 09.11.2001 |
| Inspector | E.Gudkova |

## 3.1 Mechanical Tolerances

# Mechanical Dimensions [mm]

|  | Nominal | Measured |
| --- | --- | --- |
| Transversal offset | 0.3 | 0.3 |
| Clearance | 0.05 | 0.05 |
| Vertical Straightness | 0.3 | 0.3 |
| Horizontal Straightness | 0.15 | 0.17 |
| Twist | 0.75 mrad | 0.73 mrad |

## 3.2 leak and water flow tests

# Flow Test

|  | Unit | Nominal | Measured |
| --- | --- | --- | --- |
| Pressure drop | Bar | 5 | 5 |
| Total flow | L/min | 7.9 | 8.0 |

# Pressure Test

|  | Nominal | Measured |
| --- | --- | --- |
| Pressure (bar) | 60 | 60 |
| Time (min) | 60 | 60 |
| Evidence of Leaks | No |

## 3.3 Electrical Resistance tests

# Electrical resistance at 20°C

|  | Nominal | Measured |
| --- | --- | --- |
| Temperature (°C) | 20 | 21 |
| Resistance (mohm) | 10.9 | 10.8 |

Electrical resistance between coil and ground

|  | Nominal | Measured |
| --- | --- | --- |
| Voltage (kV d.c.) | 2.0 | 2.0 |
| Temperature (°C) | 20 | 21 |
| Resistance (Ω | Min. 108 | (108 Ω |

## 3.4 Test run

# Cooling performance during test run at nominal current

|  | Checked Cutoff Cut-in |
| --- | --- |
| Functioning of thermal switch 1 | 63 0 C / 44 0 C |
| Functioning of thermal switch 2 | 65 0 C / 45 0 C |
| Functioning of thermal switch 3 | 64 0 C / 43 0 C |
| Functioning of thermal switch 4 | 63 0 C / 45 0 C |

# 4. Magnetic Measurements

# Technical Details

| Serial Number | LHC-MSIA-01 |
| --- | --- |
| Date of Inspection | 20.03.2002 |
| Inspector | A.Ivanenko |

## 4.1 System of axes MSI

## 4.2 Magnetisation curve

Flux density ( inside the magnet and integrated flux density (stretched wire method) over the total length (Bdl as a function of the excitation current.

| EXCITING CURRENT, Ampere | INTEGRAL, VOLT | B(L, Tm | Magnetic Length L, m | NMR, T |
| --- | --- | --- | --- | --- |
| 50,00 | 0,01642 | 0,14519 | 3,58128 | 0,04054 |
| 100,00 | 0,03301 | 0,29187 | 3,66675 | 0,07960 |
| 150,00 | 0,04965 | 0,43900 | 3,66386 | 0,11982 |
| 200,00 | 0,06676 | 0,59029 | 3,68447 | 0,16021 |
| 250,00 | 0,08381 | 0,74105 | 3,71079 | 0,19970 |
| 300,00 | 0,10063 | 0,88977 | 3,71355 | 0,23960 |
| 350,00 | 0,11741 | 1,03814 | 3,70763 | 0,28000 |
| 400,00 | 0,13418 | 1,18642 | 3,71626 | 0,31925 |
| 450,00 | 0,15085 | 1,33381 | 3,71597 | 0,35894 |
| 500,00 | 0,16765 | 1,48236 | 3,71629 | 0,39888 |
| 550,00 | 0,18442 | 1,63064 | 3,71731 | 0,43866 |
| 600,00 | 0,20118 | 1,77883 | 3,72186 | 0,47794 |
| 650,00 | 0,21785 | 1,92622 | 3,71980 | 0,51783 |
| 700,00 | 0,23443 | 2,07282 | 3,72207 | 0,55690 |
| 750,00 | 0,25119 | 2,22101 | 3,72635 | 0,59603 |
| 800,00 | 0,26782 | 2,36806 | 3,72529 | 0,63567 |
| 850,00 | 0,28452 | 2,51572 | 3,72953 | 0,67454 |
| 900,00 | 0,30078 | 2,65949 | 3,72665 | 0,71364 |
| 950,00 | 0,31722 | 2,80485 | 3,72579 | 0,75282 |
| 1000,00 | 0,33378 | 2,95127 | 3,72659 | 0,79195 |
| 1050,00 | 0,35036 | 3,09787 | 3,73282 | 0,82990 |
| 1100,00 | 0,36652 | 3,24076 | 3,73527 | 0,86761 |
| 1150,00 | 0,38212 | 3,37869 | 3,73443 | 0,90474 |

## 4.3 Fied map in the gap

|  |  | 5 mm towards pole |  | Current 950 A |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Int (B dl) [Tm] |  |  |  |  |  |  |
|  | End | Center | All | Homogeneity [%] | B_ave_center [T] | l_eff [m] |
| X [mm] |  |  |  |  |  |  |  |  |
| 150 | 0,16030 | 2,65820 | 2,81850 | 0,69059 |  | 0,75960 |  | 3,71051 |
| 145 | 0,16069 | 2,63401 | 2,79470 | -0,15974 |  | 0,75269 |  | 3,71295 |
| 140 | 0,16146 | 2,63359 | 2,79505 | -0,14720 |  | 0,75256 |  | 3,71407 |
| 135 | 0,16213 | 2,63516 | 2,79729 | -0,06718 |  | 0,75301 |  | 3,71481 |
| 130 | 0,16268 | 2,63749 | 2,80017 | 0,03568 |  | 0,75367 |  | 3,71540 |
| 125 | 0,16321 | 2,63854 | 2,80175 | 0,09201 |  | 0,75397 |  | 3,71598 |
| 115 | 0,16422 | 2,63894 | 2,80316 | 0,14246 |  | 0,75409 |  | 3,71728 |
| 105 | 0,16502 | 2,63869 | 2,80371 | 0,16234 |  | 0,75401 |  | 3,71840 |
| 95 | 0,16552 | 2,63767 | 2,80319 | 0,14367 |  | 0,75373 |  | 3,71908 |
| 85 | 0,16573 | 2,63647 | 2,80220 | 0,10835 |  | 0,75338 |  | 3,71951 |
| 75 | 0,16617 | 2,63609 | 2,80226 | 0,11035 |  | 0,75327 |  | 3,72014 |
| 65 | 0,16643 | 2,63584 | 2,80227 | 0,11091 |  | 0,75321 |  | 3,72046 |
| 55 | 0,16656 | 2,63499 | 2,80155 | 0,08488 |  | 0,75295 |  | 3,72076 |
| 45 | 0,16638 | 2,63454 | 2,80092 | 0,06258 |  | 0,75283 |  | 3,72051 |
| 35 | 0,16649 | 2,63401 | 2,80051 | 0,04780 |  | 0,75267 |  | 3,72078 |
| 25 | 0,16642 | 2,63397 | 2,80038 | 0,04337 |  | 0,75267 |  | 3,72062 |
| 20 | 0,16635 | 2,63402 | 2,80037 | 0,04284 |  | 0,75267 |  | 3,72056 |
| 15 | 0,16615 | 2,63401 | 2,80016 | 0,03552 |  | 0,75268 |  | 3,72027 |
| 10 | 0,16593 | 2,63411 | 2,80003 | 0,03073 |  | 0,75271 |  | 3,71993 |
| 5 | 0,16563 | 2,63401 | 2,79964 | 0,01668 |  | 0,75268 |  | 3,71957 |
| 0 | 0,16544 | 2,63396 | 2,79940 | 0,00816 |  | 0,75266 |  | 3,71934 |
| -5 | 0,16523 | 2,63377 | 2,79900 | -0,00610 |  | 0,75260 |  | 3,71910 |
| -10 | 0,16488 | 2,63363 | 2,79851 | -0,02376 |  | 0,75257 |  | 3,71859 |
| -15 | 0,16453 | 2,63377 | 2,79830 | -0,03115 |  | 0,75261 |  | 3,71813 |
| -20 | 0,16408 | 2,63420 | 2,79828 | -0,03174 |  | 0,75275 |  | 3,71742 |
| -25 | 0,16342 | 2,63197 | 2,79539 | -0,13517 |  | 0,75209 |  | 3,71681 |
| -30 | 0,16237 | 2,62790 | 2,79027 | -0,31791 |  | 0,75092 |  | 3,71580 |
| -35 | 0,16158 | 2,63040 | 2,79199 | -0,25667 |  | 0,75165 |  | 3,71445 |
| -40 | 0,16149 | 2,65054 | 2,81203 | 0,45953 |  | 0,75745 |  | 3,71250 |

|  |  | mid-plane |  | Current 950 A |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Int (B dl) [Tm] |  |  |  |  |  |  |
|  | End | Center | All | Homogeneity [%] | B_ave_center [T] | l_eff [m] |
| X [mm] |  |  |  |  |  |  |  |  |
| 150 | 0,15871 | 2,67609 | 2,83480 | 1,27275 |  | 0,76474 |  | 3,70690 |
| 145 | 0,15970 | 2,65065 | 2,81036 | 0,39959 |  | 0,75745 |  | 3,71027 |
| 140 | 0,16065 | 2,64226 | 2,80291 | 0,13351 |  | 0,75506 |  | 3,71217 |
| 135 | 0,16166 | 2,64009 | 2,80176 | 0,09237 |  | 0,75442 |  | 3,71378 |
| 130 | 0,16233 | 2,63997 | 2,80230 | 0,11181 |  | 0,75440 |  | 3,71460 |
| 125 | 0,16299 | 2,63966 | 2,80265 | 0,12446 |  | 0,75431 |  | 3,71553 |
| 115 | 0,16406 | 2,63948 | 2,80354 | 0,15623 |  | 0,75426 |  | 3,71696 |
| 105 | 0,16490 | 2,63865 | 2,80355 | 0,15655 |  | 0,75402 |  | 3,71812 |
| 95 | 0,16537 | 2,63791 | 2,80328 | 0,14685 |  | 0,75380 |  | 3,71888 |
| 85 | 0,16577 | 2,63702 | 2,80279 | 0,12934 |  | 0,75356 |  | 3,71939 |
| 75 | 0,16613 | 2,63661 | 2,80275 | 0,12775 |  | 0,75345 |  | 3,71990 |
| 65 | 0,16629 | 2,63604 | 2,80233 | 0,11290 |  | 0,75331 |  | 3,72005 |
| 55 | 0,16649 | 2,63514 | 2,80163 | 0,08780 |  | 0,75303 |  | 3,72050 |
| 45 | 0,16643 | 2,63442 | 2,80085 | 0,06005 |  | 0,75281 |  | 3,72054 |
| 35 | 0,16639 | 2,63382 | 2,80022 | 0,03734 |  | 0,75263 |  | 3,72057 |
| 25 | 0,16625 | 2,63365 | 2,79990 | 0,02602 |  | 0,75258 |  | 3,72038 |
| 20 | 0,16611 | 2,63354 | 2,79966 | 0,01736 |  | 0,75256 |  | 3,72017 |
| 15 | 0,16596 | 2,63351 | 2,79947 | 0,01086 |  | 0,75255 |  | 3,71998 |
| 10 | 0,16575 | 2,63369 | 2,79945 | 0,00987 |  | 0,75260 |  | 3,71970 |
| 5 | 0,16553 | 2,63377 | 2,79930 | 0,00469 |  | 0,75263 |  | 3,71937 |
| 0 | 0,16521 | 2,63396 | 2,79917 | 0,00000 |  | 0,75268 |  | 3,71892 |
| -5 | 0,16494 | 2,63402 | 2,79896 | -0,00762 |  | 0,75270 |  | 3,71856 |
| -10 | 0,16455 | 2,63420 | 2,79876 | -0,01482 |  | 0,75274 |  | 3,71810 |
| -15 | 0,16409 | 2,63460 | 2,79868 | -0,01739 |  | 0,75285 |  | 3,71744 |
| -20 | 0,16352 | 2,63517 | 2,79869 | -0,01716 |  | 0,75299 |  | 3,71677 |
| -25 | 0,16283 | 2,63588 | 2,79870 | -0,01665 |  | 0,75321 |  | 3,71572 |
| -30 | 0,16185 | 2,63817 | 2,80003 | 0,03058 |  | 0,75389 |  | 3,71411 |
| -35 | 0,16104 | 2,64801 | 2,80905 | 0,35284 |  | 0,75669 |  | 3,71230 |
| -40 | 0,16057 | 2,67286 | 2,83343 | 1,22385 |  | 0,76377 |  | 3,70978 |

|  |  | 5 mm towards septum | Current 950 A |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Int (B dl) [Tm] |  |  |  |  |  |  |
|  | End | Center | All | Homogeneity [%] | B_ave_center [T] | l_eff [m] |
| X [mm] |  |  |  |  |  |  |  |  |
| 150 | 0,15755 | 2,66382 | 2,82137 | 0,79308 |  | 0,76121 |  | 3,70642 |
| 145 | 0,15941 | 2,65396 | 2,81337 | 0,50713 |  | 0,75838 |  | 3,70971 |
| 140 | 0,16069 | 2,64730 | 2,80799 | 0,31510 |  | 0,75647 |  | 3,71197 |
| 135 | 0,16180 | 2,64407 | 2,80587 | 0,23936 |  | 0,75554 |  | 3,71371 |
| 130 | 0,16272 | 2,64310 | 2,80582 | 0,23754 |  | 0,75526 |  | 3,71502 |
| 125 | 0,16339 | 2,64146 | 2,80485 | 0,20276 |  | 0,75480 |  | 3,71603 |
| 115 | 0,16448 | 2,64045 | 2,80493 | 0,20584 |  | 0,75451 |  | 3,71756 |
| 105 | 0,16528 | 2,63964 | 2,80492 | 0,20529 |  | 0,75427 |  | 3,71869 |
| 95 | 0,16574 | 2,63889 | 2,80462 | 0,19484 |  | 0,75406 |  | 3,71935 |
| 85 | 0,16617 | 2,63827 | 2,80444 | 0,18814 |  | 0,75389 |  | 3,71994 |
| 75 | 0,16644 | 2,63738 | 2,80382 | 0,16597 |  | 0,75363 |  | 3,72041 |
| 65 | 0,16657 | 2,63653 | 2,80310 | 0,14036 |  | 0,75339 |  | 3,72066 |
| 55 | 0,16669 | 2,63595 | 2,80264 | 0,12380 |  | 0,75322 |  | 3,72088 |
| 45 | 0,16648 | 2,63503 | 2,80151 | 0,08372 |  | 0,75295 |  | 3,72070 |
| 35 | 0,16648 | 2,63448 | 2,80096 | 0,06393 |  | 0,75279 |  | 3,72076 |
| 25 | 0,16633 | 2,63402 | 2,80034 | 0,04193 |  | 0,75267 |  | 3,72056 |
| 20 | 0,16616 | 2,63383 | 2,79999 | 0,02915 |  | 0,75261 |  | 3,72037 |
| 15 | 0,16595 | 2,63371 | 2,79966 | 0,01766 |  | 0,75258 |  | 3,72011 |
| 10 | 0,16572 | 2,63366 | 2,79938 | 0,00766 |  | 0,75256 |  | 3,71982 |
| 5 | 0,16549 | 2,63376 | 2,79925 | 0,00297 |  | 0,75259 |  | 3,71951 |
| 0 | 0,16508 | 2,63362 | 2,79870 | -0,01685 |  | 0,75255 |  | 3,71896 |
| -5 | 0,16475 | 2,63360 | 2,79834 | -0,02959 |  | 0,75254 |  | 3,71852 |
| -10 | 0,16425 | 2,63380 | 2,79805 | -0,03998 |  | 0,75259 |  | 3,71790 |
| -15 | 0,16375 | 2,63479 | 2,79854 | -0,02249 |  | 0,75287 |  | 3,71717 |
| -20 | 0,16310 | 2,63678 | 2,79988 | 0,02522 |  | 0,75342 |  | 3,71621 |
| -25 | 0,16241 | 2,63897 | 2,80138 | 0,07884 |  | 0,75405 |  | 3,71510 |
| -30 | 0,16131 | 2,64219 | 2,80349 | 0,15440 |  | 0,75497 |  | 3,71340 |
| -35 | 0,16022 | 2,64932 | 2,80955 | 0,37069 |  | 0,75701 |  | 3,71138 |
| -40 | 0,15894 | 2,65958 | 2,81853 | 0,69148 |  | 0,75995 |  | 3,70881 |

## 4.4 field in the holes

1. Hall probe (radial),        R = 19 mm (Initial angular position equal     00

2. Hall probe (tangential)   R = 19 mm ( Initial angular position equal 2700

| S E P T U M H O L E | B E A M H O L E |
| --- | --- |

| Position degree of circle | Hall probe (radial), Gauss | Hall probe (tangential), Gauss | Absolute Value, Gauss |  | Hall probe (radial), Gauss | Hall probe (tangential), Gauss | Absolute Value, Gauss |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 4,6 | -1,5 | 4,8 |  | -3 | -2,1 | 3,7 |
| 30 | 3,9 | -2,7 | 4,7 |  | -3,6 | -0,9 | 3,7 |
| 60 | 3 | -4,3 | 5,2 |  | -4 | 0,6 | 4,0 |
| 90 | 1,3 | -6,3 | 6,4 |  | -3,8 | 3 | 4,8 |
| 120 | -2,4 | -7,8 | 8,2 |  | -1,8 | 5,3 | 5,6 |
| 150 | -8,5 | -6,1 | 10,5 |  | 2,3 | 5,5 | 6,0 |
| 180 | -12,7 | 1,6 | 12,8 |  | 5,1 | 2,9 | 5,9 |
| 210 | -8 | 10,2 | 13,0 |  | 5,4 | 0,1 | 5,4 |
| 240 | 2,5 | 10,1 | 10,4 |  | 4,5 | -2,4 | 5,1 |
| 270 | 6 | 5,5 | 8,1 |  | 2,5 | -4,2 | 4,9 |
| 300 | 6 | 2,1 | 6,4 |  | -0,3 | -4,6 | 4,6 |
| 330 | 5,2 | 0,1 | 5,2 |  | -2,2 | -3,4 | 4,0 |
| 360 | 4,6 | -1,5 | 4,8 |  | -3 | -2,1 | 3,7 |